POCKET ADDIO

12/10 BITS POCKET A/D CONVERTER

POCKET A/D TERMINAL BOARD

OPERATION MANUAL

CONTENTS:

CHAPTER 1. Introduction

CHAPTER 2. Hardware Installation

CHAPTER 3. Hardware Configuration

CHAPTER 4. Software Programming

APPENDICES:

a. Warranty Information

b. Data Sheet

Rev. Date: 1996 JAN. 22

CHAPTER 1     INTRODUCTION

The  pocket A/D is an analog to digital converter,  which  is

used  to  connect  to print port (centronic)  of  PC/AT,  PC/386,

PC/486, Pentium, or compatible computer. The pocket A/D is a high

precision data conversion system, it contains eleven 12 ( or  10)

bits analog to digital channels.

The  pocket A/D need not occupy any slot of your PC  system,

it  is  suitable for any PC which contains parallel  print  port,

especially, it is more useful for laptops, notebooks.

The features of the pocket A/D are:

* Support eleven 12 (or 10) bits channels.

* Voltage range : 0 to 5 V.

* Conversion time : 500 us.

* Current consumption : 60 mA.

* Linearity error : 1 LSB.

* Power adapter : DC 9 V, 500 mA.

The package contains:

* Pocket A/D.

* Power adapter.

* Expansion cable.

* User's manual.

* Pocket A/D terminal board (optional).

CHAPTER 2      HARDWARE INSTALLATION

Your pocket A/D is designed to be inserted in parallel print

port  of  your computer, please follow the steps  listed  in  the

following.

1.  Turn  off  all power to your  computer  and  all  peripheral

devices before installing your pocket A/D.

2.  Connect   expansion  cable to parallel  printer  port,  then

insert pocket A/D to the connector of expansion cable.

3. Plug in power adapter pocket A/D.

4. Turn on the computer and power adapter.

CHAPTER 3      HARDWARE CONFIGURATIONS

Before  you  use  the pocket A/D, you must  ensure  that  the

printer  port address and identifier address are  set  correctly.

The  proper  settings  for the pocket A/D are  described  in  the

following pages:

3.1 I/O Port address

The A/D signals are transfered via parallel printer port,  so

that  the   I/O port address of pocket A/D are &H3BC,  &H378,  or

&H278 dependenting on your printer port setting.

3.2 LED Indicator

The  LED1 is used to indicate power active and LED2 is  used

to indicate signal access.

3.3 Identifier Address

Each  expansion  cable  contains  4  connectors  to  connect

maximum 4 pocket A/D or pocket DIO (you can mix it together), the

identifier address is used to identify the address of pocket A/D.

User  may  use  DIP switch to set identifier  address,  and  each

pocket A/D must be set to different identifier address.

Observe  figure above, to set the switch ON means 0 and  set

the switch OFF means 1. The figure above is set to 5.

3.4 Pin Assignment of Pocket A/D

Pin     Function       Pin    Function

1      GND            14     +5V

2      CH0            15     GND

3      CH1            16     GND

4      CH2            17     GND

5      CH3            18     GND

6      CH4            19     GND

7      CH5            20     GND

8      CH6            21     GND

9      CH7            22     GND

10      CH8            23     GND

11      CH9            24     GND

12      CH10           25     +5V

13      GND

CHAPTER 4   SOFTWARE PROGRAMMING

In  the  distribution diskette, we provide  C,  PASCAL,  and

BASIC  examples to show how to write software to get A/D  sigals.

It is easy to call getad() function to get converted the  digital

signal  from  C and PASCAL. For BASIC language, the  function  is

located at line number 1000.

1. Programming Guide

a. Initial

void adinit(unsigned int addr);

addr    : printer port address.

b. Get A/D signal

int getad(int box, int channel);

box     : identifier address.

channel : A/D channel number.

return  : digital number.

c. Idendify bits resolution

Modify RESOLUTION parameter, it will  be

12 or 10.

2. Examples

The examples of C, PASCAL, and BASIC are

shown in the following pages.

C PROGRAM

/* Pocket A/D Converter: Sample Program */

#include <stdio.h>

#include <dos.h>

#include <conio.h>

#include <time.h>

#include <ctype.h>

/* for test only */

#define TEST

/* 12 (or 10) bits resolution */

#define RESOLUTION  12

#define LPT1 0x3BC

#define LPT2 0x378

#define LPT3 0x278

/* I/O ports for printer interface */

#define DATA(x)         (x + 0)

#define STATUS(x)       (x + 1)

#define CONTROL(x)      (x + 2)

/* definition of each data bit */

#define D7		0x80

#define D6		0x40

#define D5		0x20

#define D4		0x10

#define D3		0x08

#define D2		0x04

#define D1		0x02

#define D0		0x01

/* definition of status port */

#define BUSY	    	0x80

#define ACK         	0x40

#define PAPEROUT    	0x20

#define SELECTED    	0x10

#define ERROR       	0x08

/* definition of control port */

#define IRQENABLE	0x10

#define SELECT      	0x08

#define RESET       	0x04

#define LF          	0x02

#define STROBE      	0x01

/* pin assignments for Pocket A/D converter */

/* ADDR : box ID                            */

/* CLK  : clock for serial data in and out  */

/* DI   : data in (from host to box)        */

/* DO   : data out (from box to host)       */

/* EOC  : end of coversion                  */

#define	ADDR        	(D7 | D6 | D5 | D4)

#define CLK		D1

#define DI		D2

#define DO		SELECTED

#define EOC		ERROR

static unsigned char port_value = 0;

static unsigned int port_addr = LPT1;

/* udelay : delay routine */

/* n      : loop count    */

static void

udelay (int n)

{

int i;

for (i = 0; i < n; i++);

}

/* adinit : initial A/D converter       */

/* addr   : I/O address of printer port */

void

adinit(unsigned int addr)

{

port_addr = addr;

}

/* getad2  : kernel of A/D conversion routine */

/* next_ch : the channel which data           */

/*           will be retrived next            */

/* return  : the A/D value                    */

static int

getad2(int next_ch)

{

int i, ad;

unsigned char x;

clock_t ts;

/* wait for EOC (end of conversion) */

/* timeout = 500 ms                 */

ts = clock();

do {

if (inportb (STATUS(port_addr)) & EOC)

break;

} while (clock() - ts > 9);

if (!(inportb (STATUS(port_addr)) & EOC))

return -1;

/* start retriving data and set next channel */

ad = 0;

for (i = 0; i < RESOLUTION; i++) {

switch (i) {

case 0:

case 1:

case 2:

case 3:

x = port_value | ((next_ch & (0x08 >> i))? DI : 0);

break;

default:

x = port_value;

break;

}

outportb (DATA(port_addr), x);

ad <<= 1;

if (inportb (STATUS(port_addr)) & DO)

ad++;

outportb (DATA(port_addr), x | CLK);

udelay (1);

outportb (DATA(port_addr), x);

udelay (1);

}

return ad;

}

/* getad   : get A/D value */

/* box     : box ID        */

/* channel : A/D channel   */

int

getad(int box, int channel)

{

static int prev_channel = 0;

static int prev_box = -1;

int newflag = 0;

if (box != prev_box) {

prev_box = box;

port_value = box << 4;

outportb (DATA(port_addr), port_value);

udelay (100);

newflag = 1;

}

if (channel != prev_channel) {

prev_channel = channel;

newflag = 1;

}

if (newflag)

getad2 (channel);

return getad2 (channel);

}

#ifdef TEST

main ()

{

int lpt = 1;

int addr = LPT1;

int box[8] = {0, 1, 2, 3, 4, 5, 6, 7};

int boxno = 1;

int bi, channel, i;

unsigned long loopcount = 0;

int c;

int runstate = 0;

clrscr ();

_setcursortype (_NOCURSOR);

gotoxy (25, 2);

cputs ("Pocket A/D Converter Test Program V1.0");

gotoxy (1, 25);

cputs ("   [L]: LPT   [1-8]: BOXID   [G]: GO   [S]: STOP   [Q]: QUIT  [0]: BOXNO ");

for (channel = 0; channel < 11; channel++) {

gotoxy (10, 10 + channel);

cprintf ("%2d", channel + 1);

}

gotoxy (10, 21);

cputs ("Ref");

gotoxy (10, 22);

cputs ("Ref-");

gotoxy (10, 23);

cputs ("Ref+");

gotoxy (10, 4);

cputs ("Loop Count :");

gotoxy (40, 4);

cputs ("LPT :");

gotoxy (35, 6);

cputs ("BOX ID");

gotoxy (7, 9);

cputs ("Channel");

adinit (addr);

do {

gotoxy (23, 4);

cprintf ("%-10ld", loopcount);

gotoxy (46, 4);

cprintf ("LPT%d (%3X)", lpt, addr);

for (bi = 0; bi < boxno; bi++) {

gotoxy (20 + 6 * bi, 7);

cprintf ("BOX%d  ", bi+1);

gotoxy (20 + 6 * bi, 8);

cprintf ("%3d   ", box[bi]);

}

if (runstate == 1) {

loopcount++;

for (bi = 0; bi < boxno; bi++) {

for (channel = 0; channel < 14; channel++) {

gotoxy (19 + bi * 6, 10 + channel);

cprintf ("%5d", getad(box[bi], channel));

}

}

}

while (kbhit()) {

c = toupper(getch());

switch (c) {

case 'L':

if (++lpt > 3)

lpt = 1;

switch (lpt) {

case 1:

addr = LPT1;

break;

case 2:

addr = LPT2;

break;

case 3:

addr = LPT3;

break;

}

adinit (addr);

break;

case '1':

case '2':

case '3':

case '4':

case '5':

case '6':

case '7':

case '8':

i = c - '1';

if (++box[i] > 15)

box[i] = 0;

break;

case 13:

case 'G':

runstate = 1;

break;

case 'S':

runstate = 0;

break;

case 27:

if (runstate == 0)

runstate = -1;

else

runstate = 0;

break;

case 'Q':

runstate = -1;

break;

case '0':

if (++boxno > 8) {

boxno = 1;

for (i = 7; i < 24; i++) {

gotoxy (19, i);

clreol ();

}

}

break;

} /* switch */

} /* if */

} while (runstate != -1);

clrscr ();

_setcursortype (_NORMALCURSOR);

return 0;

}

#endif

PASCAL PROGRAM

{ Pocket A/D Converter: Sample Program }

uses

Crt,

Dos;

const

{ 12 (or 10) bits resolution }

RESOLUTION = 12;

LPT1 = $3BC;

LPT2 = $378;

LPT3 = $278;

{ I/O ports for printer interface }

DATA    = 0;

STATUS  = 1;

CONTROL = 2;

{ definition of each data bit }

D7 = $80;

D6 = $40;

D5 = $20;

D4 = $10;

D3 = $08;

D2 = $04;

D1 = $02;

D0 = $01;

{ definition of status port }

BUSY     = $80;

ACK      = $40;

PAPEROUT = $20;

SELECTED = $10;

ERROR    = $08;

{ definition of control port }

IRQENABLE = $10;

SELECT    = $08;

RESET     = $04;

LF        = $02;

STROBE    = $01;

{ pin assignments for Pocket A/D converter     }

{ ADDR     : box ID                            }

{ CLK      : clock for serial data in and out  }

{ DATAIN   : data in (from host to box)        }

{ DATAOUTO : data out (from box to host)       }

{ EOC      : end of coversion                  }

IDMASK  = D7 or D6 or D5 or D4;

CLK     = D1;

DATAIN  = D2;

DATAOUT = SELECTED;

EOC     = ERROR;

LPTTAB : array[1..3] of integer = (LPT1, LPT2, LPT3);

var

port_value : byte;

port_addr : integer;

lpt : integer;

addr : integer;

box : array[0..7] of integer;

boxno, bi, channel, i, runstate : integer;

loopcount : longint;

c : char;

{ udelay : delay routine }

{ n      : loop count    }

procedure udelay (n : integer);

var

i : integer;

begin

for i := 1 to n do;

end;

{ cursoroff : hide cursor }

procedure cursoroff;

var

r : registers;

begin

r.ah := 1;

r.ch := $20;

intr($10,r);

end;

{ cursoron : show cursor }

procedure cursoron;

var

r : registers;

begin

r.ah := 1;

r.ch := 11;

r.cl := 12;

intr($10,r);

end;

{ writehex: print hexdecimal value }

procedure writehex(n : word);

var

s : string[4];

i,z : integer;

begin

s[0] := #4;

for i := 4 downto 1 do

begin

z := n mod 16;

n := n div 16;

if z > 9 then s[i] := chr(ord('A') + z - 10)

else s[i] := chr(ord('0') + z);

end;

write(s);

end;

{ adinit : initial A/D converter       }

{ addr   : I/O address of printer port }

procedure adinit(addr : integer);

begin

port_addr := addr;

port_value := 0;

end;

{ getad2  : kernel of A/D conversion routine }

{ next_ch : the channel which data           }

{           will be retrived next            }

{ return  : the A/D value                    }

function getad2(next_ch : integer) : integer;

var

i, ad : integer;

x : byte;

begin

{ wait for EOC (end of conversion) }

{ timeout = 500 ms                 }

for i := 1 to 1000 do

if (port[STATUS + port_addr] and EOC) > 0 then break;

if (port[STATUS + port_addr] and EOC) = 0 then

getad2 := -1;

{ start retriving data and set next channel }

ad := 0;

for i := 1 to RESOLUTION do

begin

if (i <= 4) and ((next_ch and ($08 shr (i - 1))) > 0) then

x := port_value or DATAIN

else

x := port_value;

port[DATA + port_addr] := x;

ad := ad shl 1;

if (port[STATUS + port_addr] and DATAOUT) > 0 then

Inc(ad);

port[DATA + port_addr] := x or CLK;

udelay (1);

port[DATA + port_addr] := x;

udelay (1);

end;

getad2 := ad;

end;

{ getad   : get A/D value }

{ box     : box ID        }

{ channel : A/D channel   }

function getad(box : integer; channel : integer) : integer;

const

prev_channel : integer = 0;

prev_box : integer = -1;

var

newflag : integer;

begin

if box <> prev_box then

begin

prev_box := box;

port_value := box shl 4;

port[DATA + port_addr] :=  port_value;

udelay (100);

newflag := 1;

end;

if channel <> prev_channel then

begin

prev_channel := channel;

newflag := 1;

end;

if newflag > 0 then

getad2 (channel);

getad := getad2 (channel);

end;

begin

loopcount := 0;

lpt := 1;

addr := LPTTAB[lpt];

boxno := 1;

for i := 0 to 7 do

box[i] := i;

clrscr;

cursoroff;

gotoxy (25, 2);

write ('Pocket A/D Converter Test Program V1.0');

gotoxy (1, 25);

write ('   [L]: LPT   [1-8]: BOXID   [G]: GO   [S]: STOP   [Q]: QUIT  [0]: BOXNO ');

for channel := 0 to 10 do

begin

gotoxy (10, 10 + channel);

write (channel + 1:2);

end;

gotoxy (10, 21);

write ('Ref');

gotoxy (10, 22);

write ('Ref-');

gotoxy (10, 23);

write ('Ref+');

gotoxy (10, 4);

write ('Loop Count :');

gotoxy (40, 4);

write ('LPT :');

gotoxy (35, 6);

write ('BOX ID');

gotoxy (7, 9);

write ('Channel');

adinit (addr);

repeat

gotoxy (23, 4);

write (loopcount : 10);

gotoxy (46, 4);

write ('LPT', lpt, ' (');

writehex (addr);

write (')');

for bi := 0 to boxno - 1 do

begin

gotoxy (20 + 6 * bi, 7);

write ('BOX', bi+1);

gotoxy (20 + 6 * bi, 8);

write (box[bi]:3, '   ');

end;

if runstate = 1 then

begin

Inc(loopcount);

for bi := 0 to boxno - 1 do

begin

for channel := 0 to 13 do

begin

gotoxy (19 + bi * 6, 10 + channel);

write (getad(box[bi], channel):5);

end;

end;

end;

while keypressed do

begin

c := upcase(readkey);

if c = 'L' then

begin

Inc(lpt);

if lpt > 3 then lpt := 1;

addr := LPTTAB[lpt];

adinit (addr);

end

else if ('1' <= c) and (c <= '8') then

begin

i := ord(c) - ord('1');

Inc(box[i]);

if box[i] > 15 then

box[i] := 0;

end

else if (c = chr(13)) or (c = 'G') then

runstate := 1

else if (c = 'S') then

runstate := 0

else if (c = chr(27)) then

if runstate = 0 then

runstate := -1

else

runstate := 0

else if c = 'Q' then

runstate := -1

else if c = '0' then

begin

Inc(boxno);

if boxno > 8 then

begin

boxno := 1;

for i := 7 to 23 do

begin

gotoxy (19, i);

clreol;

end

end

end

end { while }

until runstate = -1;

clrscr;

cursoron;

end.

BASIC PROGRAM

100   REM

110   REM           Pocket A/D Converter: Sample Program

120   REM

130   RESOLUTION = 12

140   PV = 0

150   LPT1 = &H3BC : LPT2 = &H378 : LPT3 = &H278

160   LOOPCOUNT = 0

170   PORT = LPT1

180   BOX = 0

190   LOCATE 24, 1 : FOR I = 1 TO 25 : PRINT : NEXT

200   LOCATE 2, 20 : PRINT "Pocket A/D Converter Test Program V1.0";

210   LOCATE 24, 20 : PRINT "   [L]: LPT   [B]: BOXID   [ESC]: QUIT ";

220   LOCATE 7, 28 : PRINT "CHANNEL";

230   LOCATE 7, 45 : PRINT "VALUE";

240   FOR CHANNEL = 0 TO 10

250     LOCATE 8 + CHANNEL, 30 : PRINT USING "##"; CHANNEL + 1;

260     GOSUB 1000

270     LOCATE 8 + CHANNEL, 45 : PRINT USING "####"; AD

280   NEXT

290   LOCATE 19, 30 : PRINT "Ref";

300   CHANNEL = 11 : GOSUB 1000

310   LOCATE 19, 45 : PRINT USING "####"; AD

320   LOCATE 20, 30 : PRINT "Ref-";

330   CHANNEL = 12 : GOSUB 1000

340   LOCATE 20, 45 : PRINT USING "####"; AD

350   LOCATE 21, 30 : PRINT "Ref+";

360   CHANNEL = 13 : GOSUB 1000

370   LOCATE 21, 45 : PRINT USING "####"; AD

380   LOCATE 4, 55

390   IF PORT = LPT1 THEN PRINT "LPT1";

400   IF PORT = LPT2 THEN PRINT "LPT2";

410   IF PORT = LPT3 THEN PRINT "LPT3";

420   LOOPCOUNT = LOOPCOUNT + 1

430   LOCATE 4, 20 : PRINT "LOOP: " ; USING "##########"; LOOPCOUNT;

440   LOCATE 4, 40 : PRINT "BOX:"; BOX

450   A$ = INKEY$

460   IF A$ = "" THEN GOTO 240

470   IF A$ <> "l" AND A$ <> "L" THEN GOTO 520

480   IF PORT = LPT1 THEN PORT = LPT2 : GOTO 240

490   IF PORT = LPT2 THEN PORT = LPT3 : GOTO 240

500   PORT = LPT1

510   GOTO 240

520   IF A$ <> "b" AND A$ <> "B" THEN GOTO 560

530   BOX = BOX + 1

540   IF BOX > 15 THEN BOX = 0

550   GOTO 240

560   IF A$ <> CHR$(27) THEN GOTO 240

999   END

1000  REM      GET A/D VALUE

1010  REM      INPUT :

1020  REM           BOX     : BOX ID

1030  REM           CHANNEL : CHANNEL NUMBER

1040  REM      OUTPUT :

1050  REM           AD : A/D VALUE

1060  REM ===========================================

1070  NEWFLAG = 0

1080  IF BOX = PREVBOX THEN GOTO 1130

1090  PREVBOX = BOX

1100  PV = BOX * 16

1110  OUT PORT, PV

1120  NEWFLAG = 1

1130  IF CHANNEL = PREVCHANNEL THEN GOTO 1170

1140  PREVCHANNEL = CHANNEL

1150  NEWFLAG = 1

1160  IF NEWFLAG = 1 THEN GOSUB 1500

1170  GOSUB 1500

1180  RETURN

1500  FOR I = 1 TO 10

1510    X = INP(PORT + 1) AND 8

1520    IF X > 0 THEN GOTO 1550

1530  NEXT

1540  AD = -1 : RETURN

1550  AD = 0

1560  FOR I = 1 TO RESOLUTION

1570    IF I <= 4 THEN X = 2 ^ (4 - I) ELSE X = 0

1580    X = CHANNEL AND X

1590    IF X > 0 THEN Y = 4 ELSE Y = 0

1600    OUT PORT, PV + Y

1610    AD = AD * 2

1620    X = INP (PORT + 1) AND &H10

1630    IF X > 0 THEN AD = AD + 1

1640    OUT PORT, PV + Y + 2

1650    OUT PORT, PV + Y

1660  NEXT

1670  RETURN

APPENDIX  A  WARRANTY INFORMATION

A.1 Copyright

Copyright  1995,1996  DECISION COMPUTER  INTERNATIONAL  CO.,

LTD.  all rights reserved. No part of POCKET ADDIO  software  and

manual may be reproduced, transmitted, transcribed, or translated

into  any  language or computer language, in any form or  by  any

means,  electronic,  mechanical,  magnetic,  optical,   chemical,

manual,  or  otherwise, without the prior written  permission  of

DECISION COMPUTER INTERNATIONAL CO.,  LTD.

Each  piece  of  POCKET ADDIO package permits  user  to  use

POCKET  ADDIO  only on a single computer, a registered  user  may

use  the  program on a different computer, but may  not  use  the

program on more than one computer at the same time.

Corporate   licensing  agreements  allow   duplication   and

distribution  of  specific number of copies within  the  licensed

institution. Duplication of multiple copies is not allowed except

through  execution  of a licensing agreement.  Welcome  call  for

details.

A.2 Warranty Information

POCKET ADDIO warrants that for a period of one year from the

date  of  purchase (unless otherwise specified  in  the  warranty

card)  that  the  goods supplied will perform  according  to  the

specifications defined in the user manual.  Furthermore that  the

POCKET  ADDIO   product  will be supplied free  from  defects  in

materials  and workmanship and be fully functional  under  normal

usage.

In the event of the failure of a POCKET ADDIO product within

the specified warranty period, POCKET ADDIO will, at its  option,

replace or repair the item at no additional charge. This  limited

warranty  does  not cover damage resulting  from  incorrect  use,

electrical   interference,  accident,  or  modification  of   the

product.

All goods returned for warranty repair must have the  serial

number intact.  Goods without serial numbers attached will not be

covered by the warranty.

Transportation costs for goods returned must be paid by  the

purchaser.   Repaired goods will be dispatched at the expense  of

POCKET ADDIO.

To  ensure that your POCKET ADDIO product is covered by  the

warranty provisions, it is necessary that you return the Warranty

card.

Under this Limited Warranty, POCKET ADDIO's obligations will

be  limited to repair or replacement only, of goods found  to  be

defective as specified above during the warranty period.   POCKET

ADDIO is not liable to the purchaser for any damages or losses of

any  kind,  through the use of, or inability to use,  the  POCKET

ADDIO product.

POCKET   ADDIO   reserves  the  right  to   determine   what

constitutes warranty repair or replacement.

Return Authorization: It is necessary that any returned goods are

clearly  marked with an RA number that has been issued by  POCKET

ADDIO.   Goods  returned without this authorization will  not  be

attended to.

APPENDIX  B  DATA SHEET

-----------------------------------------------------------------

POCKET ADDIO

POCKET A/D TERMINAL BOARD

CONTENTS:

CHAPTER 1. Pocket A/D Terminal Board

APPENDIX:

APPENDIX A   Warranty Information

Rev. Date: 1996 JAN. 28

CHAPTER 1   POCKET A/D TERMINAL BOARD

The pocket A/D terminal board provides the standard  voltage

divider  circuit  for pocket A/D only, it  supports  11  channels

terminal  loop,  user  can  design  voltage  divider  circuit  by

himself.

The maximum input voltage for each channel is 0 to 5V DC  or

0 to 10V DC.

1. Terminal Board

The  figure  of pocket A/D terminal board is  shown  in  the

follows.

2. Hardware Configuration

Before  you use the A/D terminal board, you must  ensure  to

plug in this board to pocket A/D.

The LED in terminal board shows the power on status. When it

turns  to light, we have power to this board, otherwise  when  it

turns to dark, no power is provided.

3. Pin Assignment of T1

Pin  Description      Pin  Description

1      GND           14      +5V

2      CH0           15      GND

3      CH1           16      GND

4      CH2           17      GND

5      CH3           18      GND

6      CH4           19      GND

7      CH5           20      GND

8      CH6           21      GND

9      CH7           22      GND

10      CH8           23      GND

11      CH9           24      GND

12      CH10          25      +5V

13      GND

4. Pin Assignment of TB1

Pin   Description

1      CH0

2      GND

3      CH1

4      GND

5      CH2

6      GND

7      CH3

8      GND

9      CH4

10      GND

11      CH5

12      GND

5. Pin Assignment of TB2

Pin   Description

1      CH6

2      GND

3      CH7

4      GND

5      CH8

6      GND

7      CH9

8      GND

9      CH10

10      GND

6. The Use of the Jumper

JPx  is  used to select maximum input voltage, (where  x  is

channel  number) when JPx is open, the maximum input  voltage  is

10V  DC, when JPx is short, the maximum input voltage is  5V  DC.

JP0  corresponds to channel 0, and JP1 corresponds to channel  1,

... etc.

APPENDIX A   WARRANTY INFORMATION

A.1 Copyright

Copyright 1996 DECISION COMPUTER INTERNATIONAL CO., LTD. all

rights reserved. No part of POCKET ADDIO software and manual  may

be  reproduced, transmitted, transcribed, or translated into  any

language  or  computer  language, in any form or  by  any  means,

electronic,  mechanical, magnetic, optical, chemical, manual,  or

otherwise,  without  the  prior written  permission  of  DECISION

COMPUTER INTERNATIONAL CO.,  LTD.

Each  piece  of  POCKET ADDIO package permits  user  to  use

POCKET  ADDIO  only on a single computer, a registered  user  may

use  the  program on a different computer, but may  not  use  the

program on more than one computer at the same time.

Corporate   licensing  agreements  allow   duplication   and

distribution  of  specific number of copies within  the  licensed

institution. Duplication of multiple copies is not allowed except

through  execution  of a licensing agreement.  Welcome  call  for

details.

A.2 Warranty Information

POCKET ADDIO warrants that for a period of one year from the

date  of  purchase (unless otherwise specified  in  the  warranty

card)  that  the  goods supplied will perform  according  to  the

specifications defined in the user manual.  Furthermore that  the

POCKET  ADDIO   product  will be supplied free  from  defects  in

materials  and workmanship and be fully functional  under  normal

usage.

In the event of the failure of a POCKET ADDIO product within

the specified warranty period, POCKET ADDIO will, at its  option,

replace or repair the item at no additional charge. This  limited

warranty  does  not cover damage resulting  from  incorrect  use,

electrical   interference,  accident,  or  modification  of   the

product.

All goods returned for warranty repair must have the  serial

number intact.  Goods without serial numbers attached will not be

covered by the warranty.

Transportation costs for goods returned must be paid by  the

purchaser.   Repaired goods will be dispatched at the expense  of

POCKET ADDIO.

To  ensure that your POCKET ADDIO product is covered by  the

warranty provisions, it is necessary that you return the Warranty

card.

Under this Limited Warranty, POCKET ADDIO's obligations will

be  limited to repair or replacement only, of goods found  to  be

defective as specified above during the warranty period.   POCKET

ADDIO is not liable to the purchaser for any damages or losses of

any  kind,  through the use of, or inability to use,  the  POCKET

ADDIO product.

POCKET   ADDIO   reserves  the  right  to   determine   what

constitutes warranty repair or replacement.

Return Authorization: It is necessary that any returned goods are

clearly  marked with an RA number that has been issued by  POCKET

ADDIO.   Goods  returned without this authorization will  not  be

attended to.

######################################################################